## Slide 1: Coating and Performance tests of 5” PMTs for the Hall B LTC
Brandon Elman
Edward Kaczanowicz
Kyle Johnston
Zein-Eddine Meziani

## Slide 2: Bare and Coated PMT
After coating the face with
p-Terphenyl
Bare PMT
PMT 1
PMT 2
MgFl not yet applied

## Slide 3: A better view of the p-Terphenyl coating

## Slide 4: Flexible Setup for testing the PMTs*
Coated plates were used before the PMT was directly coated
We are in the process of improving the setup with a better UV lamp
PMT inside mu metal shield
Coated quartz silica plate
UV lamp
VME data aquisition
Pinhole for
Light to reach
PMT

## Slide 5: Bare PMT_1 spectrum using a Xenon lamp and a pinhole. a) PMT voltage at 2100 V b) Signal large but attenuated to less than 5V. Saturation could be an issue. The setup isn’t ideal in that the lamp is too powerful even with a small pinhole
Bare PMT spectrum

## Slide 6: Blank UV glass: Mean: 1284.44 Sigma: 12.5793 Coated UV Glass: Mean: 1340.96 Sigma: 11.9620 Gain: ~21%
For a start the tube face isn’t coated.
We compare measurement with a coated UV glass plate
vs an uncoated UV glass plate.
We clearly see a 21% gain due to the WLS

## Slide 7: Same comparison as previously but here instead of UV glass plates we use quartz glass

## Slide 8: Uncoated portion of PMT Mean: 1272.88 Sigma: 12.759
Coated portion of PMT Mean: 1325.45 Sigma: 12.1560 Gain: 19%

## Slide 9: PMT 2: Comparison between response of bare PMT vs coated one
Uncoated:
Mean: 1256.87
Sigma: 8.63327
Coated:
Mean: 1307.65
Sigma: 7.67433
Gain: 37%

## Slide 10: PMT 2: Comparison between response of bare PMT vs coated one
Uncoated:
Mean: 1257.57
Sigma: 8.59515
Coated:
Mean: 1307.65
Sigma: 7.67433
Gain: 36%

## Slide 11: PMT2 was wiped across the diameter thus most of the light was going through an uncoated area however some of it still excited the p-Terphenyl
Uncoated:
Mean: 1285.47
Sigma: 8.23189
Coated:
Mean: 1307.65
Sigma: 7.67433
Gain: 20%