## Slide 1
We expect LTCC box to be more leak-tight
New redesigned back-wall
Signals and HV connectors now hermetical
New redesigned patch panel
Various mounting fixtures fixed from the inside
Seams between wall sides and back-wall
Window production and installation
Same nose (major source of leaks)
Box Refurbishment

## Slide 2
Leak checks:
Using r134 + sniffer
Pressure drop test (~24h)
Results:
  (Rare) leaks found on window / wall: fixed
  Most leaks from nose: fixed
  No sniffer alarm after box is fixed, even at lower settings
In TED:
In The Hall Floor
Box was re-checked using same above procedure.
Some leaks appeared on the nose.
This indicates problems when moving the box, related to the nose structure.

## Slide 3
Leak checks On the Forward Carriage
On the forward carriage
2 Boxes were re-checked using 134.
Some leaks appeared on the gas line.
A hole in a window was patched.
Notes:
It seems that moving the box affects the nose.
FTOF hardware should be covered when moving LTCC.
Nose can only be fully accessed when box is out.

## Slide 4
DSG Analysis:
Price / week, all sectors:
$1470 / week all 5 sectors assuming sector 5 leak rate for all
$4508 assuming current leaks
Current C4F10 price: $140 / Kg
Note: this rates are extrapolated to C4F10. In reality leak location plays a major role due to the gas weight.
Sector | C4F10 / day | C4F10 Kg / Day | Price / Week
1 | 95 | 1.06 | 1038.8
2 | 32 | 0.36 | 352.8
3 | 118 | 1.32 | 1293.6
4 | 38 | 0.43 | 421.4
5 | 27 | 0.3 | 294
6 | 101 | 1.13 | 1107.4
 |  |  | 4508

## Slide 5
Plans
Forward carriage move first
Coordinate with Bob about Sector 4 removal date
Dave A: 134 sniff identify leaks. Fix if possible.
Sector 3 and 5 should be more accessible with 4 out.
There may be need to remove one sector to fully patch the leaks, but the plan is to leave sector 5 where it is.
DSG: Continuous leaks monitoring
Short Term
Coordinate with Bob, George for complete re-circulation and recovery system. Need DA authority
Long Term